## Slide 2
Chart: Humidity (RH %) vs Time (hr)
(1)
(2)

## Slide 3
Chart: Humidity Sensor Comparison at Equilibrium
| ~90% | ~80% | ~70% | ~60% | ~45% | ~35%
DHT-22 Average at Equilibrium | 89.54 | 80.9 | 71.05 | 60.47 | 46.27 | 34.56
Vernier Average at Equilibrium | 91.01 | 81.22 | 67.29 | 57.81 | 45.03 | 29.54

## Slide 4
Sensor Comparison in Differnet RH Environments | ~90% RH | ~80% RH | ~70% RH | ~60% RH | ~45% RH | ~35% RH
DHT-22 Average at Equilibrium (%) | 89.54 | 80.9 | 71.05 | 60.47 | 46.27 | 34.56
Vernier Average at Equilibrium (%) | 91.01 | 81.22 | 67.29 | 57.81 | 45.03 | 29.54
% Difference | 1.628357796 | 0.3947693 | 5.43588261 | 4.49780183 | 2.71632 | 15.663027
Mean of % Difference Across All Environments (%) |  |  |  |  |  | 5.0560263
Mean of % Difference w/o ~35% Condition (%) |  |  |  |  |  | 2.9346263

## Slide 5
Chart: Temperature (°C) vs Time (hr)

## Slide 6
Chart: Temperature Sensor Comparison at Equilibrium
| ~-18°C (Freezer) | ~4°C (Refrigerator) | ~21°C (Room Temp) | ~42°C (Heat Lamp-Low) | ~64°C (Heat Lamp- High)
Vernier Average at Equilibrium (°C) | -18.67 | 4.14 | 20.77 | 42.17 | 64.11
DHT-22 Average at Equilibrium (°C) | -17.93 | 3.96 | 21.1 | 41.58 | 58.71

## Slide 7
Sensor Comparison in Different Temperature Environments | ~-18°C (Freezer) | ~4°C (Refrigerator) | ~21°C (Room Temp) | ~42°C (Heat Lamp-Low) | ~64°C (Heat Lamp-High)
Vernier Average at Equilibrium (°C) | -18.67 | 4.14 | 20.77 | 42.17 | 64.11
DHT-22 Average at Equilibrium (°C) | -17.93 | 3.96 | 21.1 | 41.58 | 58.71
Calculated Vernier Sensor Error (±°C) | 0.146 | 0.212 | 0.263 | 0.326 | 0.392
DHT-22 Sensor Error (±°C) | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
DHT-22 σ at Equilibrium (°C) | 0.4345 | 0.07155 | 0.005522 | 0.1489 | 0.1097
Vernier σ at Equilibrium (°C) | 0.4822 | 0.06506 | 0.06589 | 0.2517 | 0.2571

## Slide 8
Chart: Wind Speed (mph) vs Time (min)

## Slide 9
Chart: Wind Speed Sensor Comparison at Equilibrium
| ~1.5 Setting | ~2 Setting | ~2.5 Setting | ~3 Setting | ~3.5 Setting | ~4 Setting | ~4.5 Setting | ~5 Setting
Vernier Average at Equilibrium | 5.451 | 9.989 | 17.63 | 22.81 | 28.31 | 30.23 | 34.86 | 36.07
Wind Sensor Rev. P Average at Equilibrium | 2.279 | 6.048 | 12.16 | 16.43 | 20.9 | 23.52 | 27.06 | 28.14

## Slide 10
Sensor Comparison in Different Wind Environments | ~1.5 Setting | ~2 Setting | ~2.5 Setting | ~3 Setting | ~3.5 Setting | ~4 Setting | ~4.5 Setting | ~5 Setting
Vernier Average at Equilibrium (mph) | 5.451 | 9.989 | 17.63 | 22.81 | 28.31 | 30.23 | 34.86 | 36.07
Wind Sensor Rev. P Average at Equilibrium (mph) | 2.279 | 6.048 | 12.16 | 16.43 | 20.9 | 23.52 | 27.06 | 28.14
Vernier Accuracy (±mph) | 0.336 | 0.336 | 0.5289 | 0.6843 | 0.8493 | 0.9069 | 1.0458 | 1.0821
Est. Wind Sensor Rev. P Accuracy (±mph) | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
Vernier σ at Equilibrium (mph) | 0.1767 | 0.1115 | 0.06781 | 0.06862 | 0.06913 | 0.09148 | 0.07702 | 0.1077
Wind Sensor Rev. P σ at Equilibrium (mph) | 0.07699 | 0.1341 | 0.09095 | 0.1375 | 0.1632 | 0.126 | 0.1925 | 0.2068
Average Vernier to Rev. P Measurement Ratio |  |  |  |  |  |  |  | 1.50621189
σ of Vernier to Wind Sensor Rev. P Ratio |  |  |  |  |  |  |  | 0.34730974

## Slide 11
Part | Cost
Arduino Giga R1 Wi-Fi | $72.82
DHT-22 | $7.99
Wind Sensor Rev. P | $39.95
RS485 (local shipping) | $46.10
Total: | $166.86

## Slide 12
Part | Cost
Arduino Uno Wi-Fi | $27.50
DHT-22 | $7.99
Wind Sensor Rev. P | $39.95
RS485 | $22.75
Total: | $98.19